## Sheet1
ALARM TABLE
Reference : https://e-vent.mit.edu/controls/list-of-alarms/
Type | Name | Case | Priority | Action | Solution | Column1
Pressure | High Pressure Exceeded PIP | Controller detect that the Peak Inspiratory Pressure (PIP) exceed maximum pressure value (40 cmH2O), which lead to risk of barotrauma (lung damage) | HIGH | Exhale Immediately and Stop the Ventilator, Pressure Warning LED ON | Re-run and decrease the pressure by decrease the volume/inhale speed
Pressure | Too Low Pressure | Controller detect that the plateau pressure drop significanly after the inspiratory phase (IPP) and indicates some leaks on the breathing circuit | HIGH | Exhale Immediately and Stop the Ventilator, Pressure Warning LED ON | Re-Check the Breathing Circuit connection
Pressure | Fighting Patient | Controller detect an oscillation on the pressure value on the inspiration phase | HIGH | Stop the Ventilator, Fighting warning LED ON | Check the patient condition
Volume and Pressure | Sporious Breath | Controller detect the patient inhale on the expiration phase, low pressure and increased volume value | MEDIUM | Continue the next phase and add the PEEP lost caused by the patient inhale case, Sporious Warning LED ON | If needed, re-adjust the parameter
Volume | Over Tidal Volume | Assisted volume mode, controller detect that the tidal volume delivered exceed the set tidal volume | MEDIUM | Stop the inhale phase, pause until the inhale phase over, and then continue to exhale phase, Volume Warning LED ON | If needed, re-adjust the parameter
Current | Over Current Fault | OverCurrent on the stepper motor, Ventilator accidentally stop, controller detects that there are no change in pressure or flow more than 10 sec while running | HIGH | Stop the Ventilator, Stepper driver LED ON | Check the Power Supply Unit or Motor Stepper Driver
Gas | Low/Over Oxygen | Controller detect that the oxygen concentration delivered to the patient is lower/higher than the set threshold | LOW | Oxygen LED ON | Manually adjust the oxygen flow from the oxygen tube